Sensor cable maximum

LCD display for measuring

24V DC Supply

USB output

Function keys for

We measure accurate temperature in extreme conditions

FeverSPOT EL gives digital output for communication and data logging. The 
pyrometer is equipped with a Inbuilt LCD display with red & green back lights which

shows real time temperature and counting the number of peoples. User can 
change the unit of temperature °C/°F. Temperature hold time, Set point, Reset the

storage data, Response time.

FeverSPOT EL is a rugged and compact IR Pyrometer especially designed for 
human body temperature measurement with high accuracy of 0.3°C. Its mustering

range from 0°C...50°C.

Compact, On-Line High Accurate IR Pyrometer for Body Temperature Measurement

0°C to 50°C

Accurate Sensors Technologies

length 15 mtr.

temperature and People Counting

Wall Mounting holes

parameter setting

FeverSPOT EL

Accurate Sensors Technologies

Function key to select

different functions

(Analog sub-range adjustable)

Protection Class

IP65

Features

Ÿ  Human Body Temperature Measurement

Applications

Ÿ  Extra cable lengths

Ÿ  Mechanical and Electrical Accessories

Ÿ  RS-232/RS-485 communication card

Optional

Ÿ  Pyrometer with 3 mtr. long connection cable 
Ÿ  Relay output

Ÿ  USB cable

Ÿ  Calibration Certificate, Software & Operation Manual.

Standard Scope of Supply

Normal Temperature

Fever Condition

Technical Specifications

Model 
Temperature Range

0°C to 50°C

FeverSPOT EL

Spectral Range 
Photodetector Type

Distance to Spot Size Ratio

8 - 14 μm 
Thermopile

15:1

Emissivity (ε)

0.1 to 1.2 adjustable

Response Time

Accuracy 
Repeatability

Resolution

Digital Output

Relay Output

Operating Temp. Range

Storage Temp. Range

Adjustable Parameters and 
Features via Software

100 msec adjustable upto 10 sec

0.3°C 
0.2°C

0.1°C

USB 2.0, RS-232/RS-485 interface card (Optional) 
*At a time only one digital output possible

Relay Output with hysteresis 60V DC/42V AC 
RMS,0.4A

0°C…….50°C

-20°C…70°C

Emissivity, Response Time, Unit Of 
Temperature(°C/°F), Relay Set point, Temperature

hold time etc.

Ÿ  Temperature range from 0°C to 50°C

Ÿ  Inbuilt data storage of counts with temperature 
Ÿ  15:1 optics

Ÿ  Two back lights (green & red) status indications and a buzzer 
	for alarm indication

Ÿ  Spectral range 8...14 µm

Ÿ  100 ms response time adjustable upto 10 sec 
Ÿ  USB 2.0 Digital output

Ÿ  RS-232/RS-485 Serial interface card 
Ÿ  Relay output\

Ÿ  User friendly PC software for communication

Power Supply

Power Consumption

Two options are available:

(User can select any one option) 
(1) 5V USB Adaptor

(2) 12 - 24V DC supply 
Max 1.2 watt

Housing

Operating Humidity

Weight & Dimensions

Stainless Steel

10-95%, Non-Condensing Conditions 
	200g

Dia= Ø25mm; L=103mm

FOV-15:1

0

133

2000

1000

700

400

200

67

47

27

14

6

*All Dimensions are in “mm”.

FOV-2:1

0

1000

2000

1000

700

400

200

500

350

200

100

6

Pyrometer Drawing

Software

Spot Sizes

16

5

82.5

Power supply Input 110/230V AC

Output 24 V DC, 0.7Amp.

ASTAST

Accurate Sensors Technologies

(Reference No. 9000 - 02)

20

M20 x 1.0

112.5

2 x

6 THRU

33

15

45

5

Sensing Head

Electronics

Accessories

Air Purge Unit For Sensor Head 
	(Reference No. - 8300-03)

Adjustable Mounting for Sensor Head 
	(Reference No. - 8300-04)

Converter RS-232 ↔ RS-485 
	(Reference no: 9000-03)

sold and serviced by

6260 Hawthorne Dr. Windsor, ON N8T 1J9. CANADA

Ph. (519) 948- 8371. Email: info@isolab.org